## Slide 1: CHM 5681
Raman spectroscopy
Source
hn
Sample
Detector
hn
hn

## Slide 2
Visible: 300-700 nm / 30,000-14,000 cm-1
Mid-IR: 2,500-15,000 nm / 4,000-650 cm-1
Spectroscopy
Vibration Frequency (n)
Related to:
Stiffness of the bond (k).
Atomic masses (reduced mass, m).

## Slide 3
Vibrational Spectroscopy
Infrared Spectroscopy
Raman Spectroscopy

## Slide 4
Why is the Sky blue?

## Slide 5
Wavelength Dependent Scatter

## Slide 6
Scattering depends on:
the size of the scattering particles (R)
the wavelength of light (l)
Geometric scattering (reflection): R>>l
  Rain drops (R~10-100 mm)
  All wavelengths equally scattered
  Optical effects: white clouds
Mie scattering: R~l
  Aerosols (R~0.01-1 mm)
  Red scattered better than blue
  Blue moon, blue sun
Rayleigh scattering: R<<l
  Molecules (R~0.0001-0.001 mm)
  Blues scattered better than red
  Blue sky, blue mountains, red sunsets
l
l
l
R
R
R
Wavelength Dependent Scatter

## Slide 7
This plot considers only single scattering by spheres.
Light Scattering Regime

## Slide 8
Geometric scattering (reflection): R>>l

## Slide 9
This plot considers only single scattering by spheres.
Light Scattering Regime

## Slide 10
Mie Scattering
Mie scattering: R~l
  Aerosols (R~0.01-1 mm)
  Red scattered better than blue
  Blue moon, blue sun
l
R
All directions but not isotropic.
Dependent on particle radius and refractive index.
Much more intense than Rayleigh scattering.
Not as strongly wavelength dependent (l2).
For colloids in solution it is also known as Tyndall Effect.
Small Particle
Large Particle

## Slide 11
Mie Scattering/Tyndall Effect
Opalescent glass
Dynamic Light Scattering (Mattoussi lab)

## Slide 12
Mie Scattering and Doppler Radar
Hurricane Michael in October 2018
Microwaves

## Slide 13
Light Scattering Regime
Atoms
Molecules
Nanoparticles

## Slide 14
Rayleigh Scattering
Rayleigh scattering: R<<l
  Molecules (R~0.0001-0.001 mm).
  What we observe with
  Raman spectroscopy.
l
R
Incident light causes a momentary distortion of the electrons distributed around a bond.
This induces a temporary dipole moment.
The distorted cloud relaxes back to its original state.
A photon is “re-emitted” during relaxation.
The “emitted” photon is usually the same energy as the incident photon.

## Slide 15
Rayleigh Scattering
I = scattering intensity at q
I0 = incident light intensity
q = angle between incident and scattered light
a = polarizability
l = wavelength
R = distance from the scatterer
Take homes:
Dependent on a2
Dependent on l4
400 nm scatters ~10 times as much as 700 nm
Source

## Slide 16
Rayleigh Scattering
400 nm scatters ~10 times as much as 700 nm
Source
Atmospheric composition: N2 (78%), O2 (21%), Ar (1%)
Size of N2 molecule: 0.31 nm
Size of O2 molecule: 0.29 nm
Size of Ar molecule: 0.3 nm
Only ~1 in 106 of the incident photons scatter.

## Slide 17
Rayleigh Scattering
I = scattering intensity at q
I0 = incident light intensity
q = angle between incident and scattered light
a = polarizability
l = wavelength
R = distance from the scatterer
Take homes:
Dependent on a2
Dependent on l4
“Softness” of wave function.
Larger clouds more easily polarized by an electric field.
More polarizable, the greater the induced dipole and scatter.
Polarization of a bond is symmetry dependent.
Raman spectroscopy is symmetry dependent.
a = polarizability

## Slide 18
1) Find number/symmetry of vibrational modes.
2) Assign the symmetry of known vibrations.
3) Find if a vibrational mode is IR or Raman Active.
Vibrations and Group Theory

## Slide 19
If the symmetry label of a normal mode corresponds to x, y, or z, then the fundamental transition for this normal mode will be IR active.
If the symmetry label of a normal mode corresponds to products of x, y, or z (such as x2 or yz) then the fundamental transition for this normal mode will be Raman active.
Selection Rules
Change of the polarization.
Deformation of an electron cloud.
Raman Active
IR Active
Change in dipole moment.
Polarizability has the same symmetry properties as the quadratic functions!

## Slide 20
Vibrations = A1 + B2
Stretch
Stretch
Bend
A1
If the symmetry label of a normal mode corresponds to x, y, or z, then the fundamental transition for this normal mode will be IR active.
If the symmetry label of a normal mode corresponds to products of x, y, or z (such as x2 or yz) then the fundamental transition for this normal mode will be Raman active.
IR vs Raman Active
A1 and B2 are Raman Active.

## Slide 21
Rayleigh Scattering
I = scattering intensity at q
I0 = incident light intensity
q = angle between incident and scattered light
N = number of scatterers
a = polarizability
l = wavelength
R = distance from the scatterer
Take homes:
Dependent on a2
Dependent on l4
If Raman active it will inelastically scatter incident light.
Elastic Scattering
Inelastic Scattering
hn
Sample
hn
hn
Sample
hn
Incident and scattered light the same energy.
Most scatter is elastic.
Incident and scattered light NOT the same energy.
Much more rare.

## Slide 22
1871 – Rayleigh elastic scattering described
1923 – Inelastic light scattering is predicted by A. Smekel
1928 – C.V. Raman and K.S. Krishnan see “feeble fluorescence” from neat solvents
Inelastic Scattering

## Slide 23
1871 – Rayleigh elastic scattering described
1923 – Inelastic light scattering is predicted by A. Smekel
1928 – C.V. Raman and K.S. Krishnan see “feeble fluorescence” from neat solvents
Inelastic Scattering
Hg arc lamp
Spectrum on photographic plate.
Hg arc lamp
through benzene
New “emission” lines.
Shift not dependent on incident wavelength.
Shift dependent on structure of molecules.
Gas, solids, and liquids.

## Slide 24: Raman Spectroscopy
1871 – Rayleigh elastic scattering described
1923 – Inelastic light scattering is predicted by A. Smekel
1928 – C.V. Raman and K.S. Krishnan see “feeble fluorescence” from neat solvents
1928 – Landsberg and Mandelstam see unexpected frequency shifts in scattering from quartz
1930 – C.V. Raman wins Nobel Prize in Physics
1957 – First confocal Raman spectrometer
1961 – Invention of laser makes Raman experiments easier
1976 – Time-resolved Raman
1977 – Surface-enhanced Raman scattering (SERS) is discovered
1997 – Single molecule SERS is possible
Inelastic Scattering

## Slide 25
Cyclohexane Scattering
488 nm laser
488 nm laser + notch filter
Elastic Scattering
Inelastic Scattering

## Slide 26
Raman Scattering
S0
S1
Energy
n0
n1
n2
S0
S1
Energy
n0
n1
n2
S0
S1
Energy
n0
n1
n2
Virtual state
Rayleigh
Elastic
Raman
Stokes shift
Raman
anti-Stokes shift
nex = nscatter
nex < nscatter
nex > nscatter
 = nex - nscatter
= vib

## Slide 27
S0
Energy
n0
n1
n2
S0
Energy
n0
n1
n2
S0
Energy
n0
n1
n2
Virtual state
Raman Spectra
Rayleigh
anti-Stokes
Stokes
Stokes
anti-Stokes
Rayleigh
Intensity (counts)
Energy (cm-1)

## Slide 28
Stokes
anti-Stokes
nex + nvib
Rayleigh
nex - nvib
Intensity (counts)
107 counts
E = nvib
Raman Spectra
1 count

## Slide 29
nex - nvib
Intensity (counts)
Raman Spectra
nvib
S0
Energy
nvib

## Slide 30
Excitation Energy, s (cm–1)
Intensity
11,000
13,000
15,000
17,000
19,000
21,000
Near IR
785 nm
Visible
514 nm
–∆E
+∆E
–∆E
+∆E
Stokes
Anti-Stokes
Wavelength Independent
Stokes
Anti-Stokes
S0
S1
Energy
n0
n1
n2
785 nm
514 nm

## Slide 31
nex - nvib
Intensity (counts)
107 counts
1 count
Intensity of Raman Peaks
<0.01% of incident photons scatter elastically.
<0.00000001% scatter inelastically.
Intensity/ratio dependent on:
Intensity of the source
Concentration of the molecule
Polarizability of the molecule
Excitation wavelength
  Dependent on l4
  Higher energy = decomposition

## Slide 32
Intensity of Stokes vs Anti-stokes
Stokes and anti-Stokes shift are the same.
Istokes is “always” greater than Ianti-Stokes
S0
n0
n1
n2
S0
Energy
n0
n1
n2
anti-Stokes
Stokes
Excitation probabilities are similar.
Concentration of n0 > n1.
Follows Boltzmann distribution.
n0 > n1
Nlower > Nupper
Increasing temp
increases Ianti-Stokes
decreases Istokes

## Slide 33
Increasing Signal Intensity
Non-Resonant
Resonant
Tacoma Narrows Bridge (1940)

## Slide 34
Increasing Signal Intensity
S0
S1
Energy
n0
n1
n2
Virtual state
Non-Resonant
Resonant
S0
S1
Energy
n0
n1
n2
Not absorption and emission!
Resonant transition = more polarizable
Intensity increase by:
x10 for MC/d-d transition
to
x100,000 for π-π* transition
<0.00000001% of incident photons
 is the same.

## Slide 35
Problem with Resonance Raman
<0.001%
<0.0000001%
100-0%
<0.01%
lex
Abs
Stokes Raman Scatter
Fluorescence
Stokes
Raman
Spectrum
Raman + Fluorescence
lex

## Slide 36
Stokes Raman Spectrum
Raman + Fluorescence
Avoiding Fluorescence
Use non-resonant Raman
  -Sub-Band Gap light
  -Red or nearIR laser
Temporal Gating
  -Raman is fast (<ps)
  -Emission is “slow” (> ps)
Anti-stokes Raman Spectroscopy
  -Emission is “always” stokes
Polarization
  -Emission if usually isotropic
  -Scattering is polarized
lex

## Slide 37
Raman Spectrometer
Procedure
Step 1: Prepare a sample
Step 2: ???
Step 3: Obtain spectra (Profit!)
Source
hn
Sample
Detector
hn
hn

## Slide 38
Raman Spectrometer
Detector
Source
1) Light source
2) Sample
3) Wavelength Selector
4) Detector
5) Filters, Slits, Polarizers

## Slide 39
Sources
Source
hn
Sample
Detector
hn
hn
Ideal
  -Narrow bandwidth
  -High intensity
Theoretically could be any light source + monochrometer or filter.
Almost always a laser
Wavelength (nm)
Argon Ion Laser

## Slide 40
Sources

## Slide 41
Detectors
PMT
Full spectra detection
Single l detection
Diode
CCD
Diode Array
hn
Detector
electrical signal
      high sensitivity
      high signal/noise
      constant response for λs
Unusually the same as UV-Vis/Fluorescence Spectroscopy

## Slide 42
Gas or Liquid Samples
      Cuvette
      Vial
      NMR Tube
Transparent to visible and near IR light.
Raman inactive: quartz, glass, some plastics.
Test sample holder and solvent for Raman peaks.
Unlike transmission IR, can be used with water.
-Water = weak Raman scatterer
-Biologically relevant structures

## Slide 43
Sample size ≈ laser spot size
Solid Samples
Neat Solids
Thin Films
Capillary Tube
Tissue
Little to no sample prep.
Antiquities

## Slide 44
Raman Spectrometer
Detector
Source
1) Light source
2) Sample
3) Wavelength Selector
4) Detector
5) Filters, Slits, Polarizers

## Slide 45
Wavelength Selector: Filters
Narrow Bandpass Filter
Notch Filter

## Slide 46
Wavelength Selector
Grating
Dispersive Spectrometer (visible)
Interferometer
FT Spectrometer (nearIR-IR)

## Slide 47
Wavelength Selector: Grating
Dispersive Spectrometer (visible)

## Slide 48
Wavelength Selector: Interferometer
FT Spectrometer (nearIR-IR)

## Slide 49
FT-Raman vs. Dispersive Raman
Advantages of FT
NIR lasers reduce fluorescence.
Relatively inexpensive.
High resolution.
Collects Stokes and Anti-stokes simultaneously.
Can be directly coupled to an FT-IR spectrometer.
Disadvantages of FT
Sample heating + blackbody emission.
Lower scattering (l4).
NIR absorption bands.
Slow (tens of minutes).

## Slide 50
Raman Spectrometer
Source
hn
Sample
Detector
hn
hn
Raman Measurements
  Microscope
  Polarized
  Confocal
  SERS
  Time Resolved
  Portable

## Slide 51
Raman Microscope

## Slide 52
Raman Microscope
Stage
Objective lens
Dichroic
Mirror
Camera
Edge filter
Focusing lens
Computer
Mono-
chromator
CCD
Optical fiber
LASER

## Slide 53: Raman Instrumentation
Tunable Laser System
Versatile Detection System
Raman Microscope
Narrow peaks.
Need higher resolution.

## Slide 54
Raman Microscope

## Slide 55
Raman Microscope Map

## Slide 56
Horiba JY LabRam HR800 Raman Spectrometer
In House Raman Spectrometer
Excitation:
785 nm, 633 nm, 514.5 nm,
488 nm, 458 nm & 364 nm
5x, 10x, 20x, 40x & 50x Visible/NIR objectives
CCD detector
CSL 3701
MAC Lab
(Materials Characterization)

## Slide 57
Raman Spectrometer
Source
hn
Sample
Detector
hn
hn
Raman Measurements
  Microscope
  Polarized
  Confocal
  SERS
  Time Resolved
  Portable

## Slide 58
Polarized Raman Spectroscopy

## Slide 59
Raman Spectrometer
Source
hn
Sample
Detector
hn
hn
Raman Measurements
  Microscope
  Polarized
  Confocal
  SERS
  Time Resolved
  Portable

## Slide 60
Raman Microscopy
Conventional
Confocal

## Slide 61
Raman Microscopy

## Slide 62
Confocal Raman

## Slide 63
75 m
Polyethylene
Polyethylene
nylon
z
x
Confocal Raman
Depth Profiling

## Slide 64
500
1000
1500
2000
film
adhesive
500
1000
1500
2000
Glass
Confocal Raman
Depth Profiling

## Slide 65: Confocal Aperture and Field Depth
Confocal Raman
3D Imaging

## Slide 66
Raman Spectrometer
Source
hn
Sample
Detector
hn
hn
Raman Measurements
  Microscope
  Polarized
  Confocal
  SERS
  Time Resolved
  Portable

## Slide 67
Surface Enhanced Raman Scattering (SERS)
M. Fleischmann, P. J. Hendra, and A. J. McQuillan. Raman spectra of pyridine adsorbed at a silver electrode. Chem. Phys. Lett., 1974, 26, 163.
solution
surface

## Slide 68
Surface Enhanced Raman Scattering (SERS)
Samples adsorbed on the surface of colloidal metal particles (usually silver, gold, or copper)
Enhancement of the electromagnetic field surrounding molecule near plasmon resonance.
Large local field increases the absorption cross section.
Scattered radiation is also enhanced.
Signal increase by a factor of 106 to 108.

## Slide 69
Surface Enhanced Raman Scattering (SERS)
E. C. Le Ru et al., J. Phys. Chem. C, 2007, 111, 13794

## Slide 70
Surface Enhanced Raman Scattering (SERS)
Substrate

## Slide 71
Raman Spectrometer
Source
hn
Sample
Detector
hn
hn
Raman Measurements
  Microscope
  Polarized
  Confocal
  SERS
  Time Resolved
  Portable

## Slide 72
C+
A-
C
A
C*
A
Transient Absorption Spectroscopy

## Slide 73
C+
A-
C
A
C*
A
Time-Resolved Raman

## Slide 74
Time-Resolved Raman
S0
S1
Energy
n0
n1
n2
Virtual state
Ground State Raman
Excited State Raman
Virtual state

## Slide 75
Events:
1) Raman Spectrum
2) Excitation Flash
3) Raman Spectra
Time
Raman probe
pump
probe
probe
probe
Excited Processes
Time-Resolved Raman
hn
(1)
(1)
(2)
(3)
Source
hn
Sample
Detector
hn
(3)
Detector

## Slide 76
I for xS0 molecules
Difference Spectra
hn
4 excited states/100 molecules
I for (x - y)S0 + yS1 molecules
S0
S1
E

## Slide 77
Difference Spectra
I(t) - I(0) = DI
-
=
DI at time t
I(t)
I(0)
I for
xS0
I for
(x - y)S0 + yS1
- yS0 + yS1
I(0) = intensity without laser pulse
I(t) = intensity at time t after laser pulse

## Slide 78
Time-Resolved Raman (10-9 s)
OPO: Optical Parametric Oscillator, RS: Raman Shifter, DG: Delay Generator, PB: Pellin-Broca Prism, DM: Dichroic Mirror, PD: Photodiode, L1: Focusing Lens, L2: Collection Optics, NF: Notch Filter (required for single monochromator), MONO: Double Monochromator (single and triple monochromator can also be used), CCD: Charge Coupled Device, ICCD: Intensified CCD, EMCCD: Electron Multiplying CCD, PDA: Photodiode Array, OSC: Oscilloscope, CP: Circulating Pump, and PC: Personal Computer

## Slide 79
Time-Resolved Raman (10-12 s)

## Slide 80
Time-Resolved Raman
J. Chem. Phys. 2001, 115, 11228.

## Slide 81
Raman Spectrometer
Source
hn
Sample
Detector
hn
hn
Raman Measurements
  Microscope
  Polarized
  Confocal
  SERS
  Time Resolved
  Portable

## Slide 82
Portable Raman Spectrometer

## Slide 83
Raman Spectrometer
Source
hn
Sample
Detector
hn
hn
Raman Measurements
  Microscope
  Polarized
  Confocal
  SERS
  Time Resolved
  Portable

## Slide 84
Identification of inorganic, organic, and biological samples
Identification of components of an unknown mixture
Fingerprint spectra
Identification of crystalline phases and orientations
Measure stress on solids
Analysis of solids, liquids, and gasses
Component quantification
Forensic sciences
In measurement and analysis of Atmospheric Spectra
Applications

## Slide 85
Nasdala, L., Harris, J.W. & Hofmeister, W. (2007): Micro-spectroscopy of diamond. Asia Oceania Geosciences Society, 4th Annual Meeting, Bangkok, Thailand, August, 2007.
Nasdala, L., Raman barometry of mineral inclusions in diamond crystals. Mitt. Österr. Miner. Ges. 149 (2004)
Mechanical Strain

## Slide 86
Zircon (ZrSiO₄): amorphous (blue) and crystalline (red).
Amorphous vs Crystalline

## Slide 87
Anatase and Rutile TiO2 –
identical chemical composition, different crystalline structures.
Rutile
Anatase
Polymorph Determination

## Slide 88
Composition
plastic

## Slide 89: Ivory or Plastic?
Composition
1725 cm-1 = pmma
1086 cm-1 = calcite

## Slide 90
Brown, K. L.; Clark, J. H. R., Anal. Chem. 2002, 74,3658.
“15th century” map
Composition
1920s

## Slide 91
Tissue Identification

## Slide 92
In contrast to fluorescent imaging, it does not require an emissive molecule/label.
Tissue Identification

## Slide 93
Vibrational Spectroscopy
Infrared Spectroscopy
Raman Spectroscopy
vs.

## Slide 94
Comparing Vibrational Spectroscopies
Advantages of Raman Over IR
Can use visible light sources and detectors.
Glass and quartz sample holders can be used.
Symmetric vibrations can be observed (crystalline solids)
Raman intensities proportional to laser power.
Minimal sample prep.
Raman bands are polarized (orientation information).
Confocal for spatial imaging.
Readily observe 100 to 700 cm-1 region (M-L bonds).
Water can be used as a solvent (biological native state).
Disadvantages of Raman vs. IR
Very weak signals (106 photons for 1 Raman photon)
Requires a laser.
More expensive.
More difficult (laser alignment, ambient lighting).

## Slide 95
Complementary Techniques

## Slide 96
Complementary Techniques

## Slide 97
Raman Spectroscopy End
Any Questions?